ECG LENS

# Experiment Pipeline
Defence Guide

Viva preparation • Decisions, alternatives, results, limitations

REFERENCE RUN USED IN THIS GUIDE The original complete experiment run: 44 of 44 eligible recordings processed; 100,685 heartbeat examples; selected validation ensemble = class-weighted Linear SVM + XGBoost + AdaBoost; validation macro-F1 = 0.5460.

## What this document is for

A practical oral-defence companion for explaining the experiment notebook from dataset selection through preprocessing, feature engineering, leakage-aware splitting, baseline modelling, class-imbalance handling, voting/stacking comparisons, final test evaluation and model-bundle saving.

## Core message

This was an experimental classical-ML pipeline for four-class ECG heartbeat classification. The chosen ensemble is the best by validation macro-F1 among the predefined voting combinations tested in the original complete run; it is not claimed to be universally optimal or clinically validated.

## How to use it

Read Sections 1–9 to understand the decisions and numbers.

Practise the viva answers in Sections 10–11 aloud; adapt the wording naturally.

Use the final “numbers to remember” page the night before the discussion.

Source basis: the notebook snapshot named “mit-bih-ipynb (13).ipynb” supplied in this conversation and the original complete-run outputs recorded in the conversation. This guide deliberately does not mix in the later rerun that failed to download record 202.

# 1. The project in one minute

Suggested opening answer:

60–90 SECOND OPENING “ECG Lens is a classical machine-learning experiment for classifying individual ECG heartbeats into four AAMI-style groups: N, S, V and F. I inspected the MIT-BIH annotations, excluded four paced recordings under the selected protocol, filtered the MLII lead, and extracted fixed 216-sample heartbeat windows. I combined those waveform values with eight amplitude/QRS statistics and two RR-interval features, producing 226 features per heartbeat. I then used subject-grouped train, validation and test partitions to reduce patient leakage, compared Logistic Regression, Linear SVM, Random Forest, AdaBoost and XGBoost, and explored class weighting, hard voting and stacking. The SVM + XGBoost + AdaBoost voting combination had the best validation macro-F1 among the predefined voting combinations, at 0.5460. On the untouched held-out test partition it achieved 86.32% accuracy, 56.87% balanced accuracy and 0.5161 macro-F1. The minority-class results remain limited, so I present this as an educational prototype and an experiment in handling imbalanced classification, not a clinical diagnostic tool.”

## Pipeline at a glance

| Stage | Input → operation | Output / reason |
| --- | --- | --- |
| 1. Data inspection | MIT-BIH ECG waveforms + expert annotation files | Understand signal channels, annotation symbols and label imbalance. |
| 2. Target definition | Raw beat symbols → N / S / V / F | A focused four-class classification problem; Q is excluded. |
| 3. Preprocessing | MLII → 4th-order Butterworth 0.5–40 Hz, zero-phase | Reduced baseline drift/high-frequency components while retaining visible QRS morphology. |
| 4. Segmentation | Annotation sample ± 0.2 s before / 0.4 s after | 216-value heartbeat window at 360 Hz. |
| 5. Features | 216 waveform + 8 summary/QRS + 2 RR values | 226 model inputs; RR boundary values may be missing. |
| 6. Split | StratifiedGroupKFold with subject groups preserved | Training 58,295; validation 18,869; test 23,521 examples. |
| 7. Experiments | Baselines → tuning → class weighting → boosting → voting/stacking | Select using validation macro-F1, inspect class-level performance. |
| 8. Final evaluation | Chosen vote on held-out test set, SVM in case of no clear majority | 86.32% accuracy; 56.87% balanced accuracy; 0.5161 macro-F1. |
| 9. Persistence | Save fitted estimators, imputer, schema and settings | Reusable bundle for the separate prediction notebook. |

# 2. Dataset, labels and scope

## 2.1 Why MIT-BIH?

The experiment uses the MIT-BIH Arrhythmia Database because it provides ECG signals together with expert beat annotations. Those annotations make supervised heartbeat-window extraction possible. WFDB is used to read recordings and annotation files.

## 2.2 Record selection

The database list contains 48 record identifiers; four paced recordings (102, 104, 107 and 217) were excluded under the chosen protocol.

The original reference run successfully processed all 44 eligible recordings. Recordings 201 and 202 are retained in the same subject group because they belong to the same patient.

The selected channel is MLII when available. The notebook reports MLII selected for all 44 eligible recordings.

## 2.3 Mapping original symbols to four target classes

| Target class | Original annotation symbols mapped | Interpretation used here |
| --- | --- | --- |
| N | N, L, R, e, j | Normal and related beats |
| S | A, a, J, S | Supraventricular ectopic beats |
| V | V, E | Ventricular ectopic beats |
| F | F | Fusion beats |
| Q (excluded) | /, f, Q | Paced / unclassifiable group in this experiment |

Other annotation symbols that do not belong to the mapping (such as rhythm-change and non-beat/event markers) are not turned into model targets. The experiment is a four-class beat classifier; it does not claim to classify every annotation type in MIT-BIH.

## 2.4 Why exclude Q and paced records?

The initial scope is a defined four-class task rather than a five-class task that also includes paced/unclassifiable beats. The four paced recordings are excluded according to the chosen protocol. This narrows the task and helps keep the input domain more consistent, but it also means the model cannot answer questions about the excluded Q category.

## 2.5 Original complete-run dataset size

| Class | Heartbeat examples | Approx. share |
| --- | --- | --- |
| N | 90,094 | 89.48% |
| S | 2,781 | 2.76% |
| V | 7,008 | 6.96% |
| F | 802 | 0.80% |
| Total | 100,685 | 100% |

IMPORTANT INTERPRETATION This is a highly imbalanced dataset: normal beats dominate, and fusion beats are rare. That is why raw accuracy alone is not a sufficient model-selection criterion.

# 3. Signal preprocessing and feature engineering

## 3.1 Filtering decision

| Setting | Choice | Defence rationale |
| --- | --- | --- |
| Lead | MLII, preferred when available | Keeps one consistent input channel instead of mixing leads with different morphology. |
| Filter | Butterworth band-pass | A standard smooth pass-band response for reducing slow drift and high-frequency components. |
| Cutoffs | 0.5–40 Hz | The notebook compares the raw and filtered waveform and uses this band for the experiment. |
| Order | 4th order | Defined filter complexity used consistently in the training pipeline. |
| Application | SOS + `sosfiltfilt` | Second-order sections improve numerical stability; forward-backward filtering avoids phase shift. |
| Sampling | 360 Hz | MIT-BIH sampling rate used for window and RR timing calculations. |

What to say accurately: “The filtered plots showed that the major ECG peaks remained visible and the baseline was more centred. I therefore used these settings consistently.” Do not claim that filtering was proven to improve classification by an ablation experiment; the notebook visualises and adopts the filter but does not establish a controlled raw-versus-filtered model comparison.

## 3.2 Why 216 samples per beat?

At 360 samples per second, 0.2 seconds before the annotation gives 72 samples and 0.4 seconds after gives 144 samples. Together, the 0.6-second window contains 216 values. The annotated location is at index 72 within the window. Keeping one fixed window length gives classical ML models a fixed-width feature vector.

Windows crossing the start or end of a recording are dropped if a complete window cannot be constructed.

Segments are centred on expert annotation locations during training; the separate inference pipeline uses an automatic detector, which is an additional source of possible differences.

## 3.3 The 226-feature vector

| Feature group | Count | Features / purpose |
| --- | --- | --- |
| Waveform samples | 216 | Every sample in the fixed-length filtered segment; preserves local morphology. |
| Amplitude statistics | 6 | Mean, standard deviation, minimum, maximum, peak-to-peak amplitude and RMS amplitude. |
| QRS statistics | 2 | QRS-slice peak-to-peak amplitude and QRS-slice RMS, using segment indices [54:91]. |
| Timing / rhythm context | 2 | Previous and next RR intervals in seconds. |
| Total | 226 | 216 + 8 + 2. |

## 3.4 Why combine raw samples, statistics and RR intervals?

The raw waveform values retain shape details that may be lost if the beat is represented only by summary statistics.

Amplitude and QRS statistics offer compact descriptors of waveform size and the QRS region.

RR intervals add temporal context: the timing between adjacent beats can help distinguish rhythms even when local shapes overlap.

The first beat may have no previous RR and the last may have no next RR. These boundary NaNs are expected; they are imputed using a median fitted on training data only.

LEAKAGE PREVENTION Do not fit an imputer or scaler on the entire dataset before splitting. In this experiment, the subject split comes first; model preprocessing is then fitted on training examples and applied to validation/test data.

# 4. Patient-independent split and evaluation metrics

## 4.1 Why split by subject rather than by beat?

Beats from one person share signal characteristics. A random beat-level split could place beats from the same person in both training and test sets and make the evaluation look better than performance on an unseen person. The notebook therefore uses StratifiedGroupKFold-style group preservation; records 201 and 202 are grouped as one subject.

## 4.2 Original complete-run partition

| Partition | Examples | Unique subject groups | N | S | V | F |
| --- | --- | --- | --- | --- | --- | --- |
| Train | 58,295 | 25 | 53,752 | 1,997 | 2,144 | 402 |
| Validation | 18,869 | 8 | 16,762 | 192 | 1,890 | 25 |
| Test | 23,521 | 10 | 19,580 | 592 | 2,974 | 375 |

The notebook deliberately assigned fusion-rich subject groups across the partitions: subject_208 to training, subject_223 to validation and subject_213 to testing. The intent was to make the rare F class available for model development and final evaluation without splitting an individual subject across partitions.

LIMITATION OF THIS SPLIT This is a purposeful subject-independent split, not a random draw designed to represent the population prevalence. The validation set contains only 25 F examples and the test set 375; class-wise estimates, especially for F, should be interpreted cautiously.

## 4.3 Metrics: what each one tells us

| Metric | Meaning | Why it is reported |
| --- | --- | --- |
| Accuracy | Fraction of all examples classified correctly. | Easy overall summary but can be dominated by N. |
| Balanced accuracy | Mean recall across the classes. | Shows class-level sensitivity without giving N more weight just because it is more common. |
| Macro-F1 | Mean of the per-class F1 scores, with equal class weighting. | Primary selection metric: rewards a balance of precision and recall for each class. |
| Precision | Of predictions assigned to a class, how many were correct. | Useful when false positives matter. |
| Recall | Of true examples of a class, how many were found. | Useful for identifying which classes are being missed. |
| Confusion matrix | Counts of true vs predicted classes. | Shows which classes are confused with which others. |

Macro-F1 was declared the model-selection metric because 89.48% of extracted beats are N. A classifier can obtain high accuracy by doing well on N while performing poorly on S or F. Macro-F1 prevents N prevalence from completely dominating the selection score.

## 4.4 Why use a validation set and a separate test set?

The training set fits the models and preprocessing. The validation set is used for tuning, weighting decisions and choosing the ensemble. The held-out test partition is reserved for the final evaluation after model choice. If the test set were repeatedly used to choose models, it would stop being an independent final check.

# 5. Baselines, tuning and class weighting

## 5.1 Why start with three classical models?

| Model | Role in the experiment | Preprocessing |
| --- | --- | --- |
| Logistic Regression (LR) | Simple linear baseline; fast and useful to compare a linear decision boundary. | Median imputation + standard scaling + LogisticRegression. |
| Linear SVM | Maximum-margin linear classifier; a useful strong baseline for high-dimensional features. | Median imputation + standard scaling + LinearSVC. |
| Random Forest (RF) | Nonlinear tree-ensemble baseline; can model feature interactions without scaling. | Median imputation; no standard scaling. |

## 5.2 Baseline and tuned validation results

| Model | Baseline accuracy | Baseline macro-F1 | Tuned accuracy | Tuned macro-F1 |
| --- | --- | --- | --- | --- |
| Logistic Regression | 93.33% | 0.4390 | 94.48% | 0.4470 |
| Linear SVM | 93.48% | 0.4183 | 93.37% | 0.4125 |
| Random Forest | 86.25% | 0.3619 | 95.53% | 0.4590 |

The tuned Random Forest led these unweighted/tuned baselines by macro-F1, but its score still reflects the class-imbalance challenge. It had very weak S recall and only 0.04 F recall in the tuned validation report. The lesson was not “Random Forest is useless”; it was that baseline ranking alone does not solve minority-class recognition.

## 5.3 Hyperparameter search design

Logistic Regression and Linear SVM: searched C ∈ {0.01, 0.1, 1.0, 10.0} using three-fold StratifiedGroupKFold on the training partition.

Random Forest: a bounded grid of 4 settings crossed with 3 group-aware folds: n_estimators ∈ {50, 100}; max_depth ∈ {15, 25}; min_samples_leaf = 5; max_features = sqrt.

The best recorded group-CV macro-F1 was 0.4282 for SVM at C=0.01, 0.4076 for LR at C=0.01, and 0.3453 for RF with max_depth=15, max_features=sqrt, min_samples_leaf=5, n_estimators=50.

The bounded searches are experiments, not proof that the global optimum was found. The search ranges are finite, and the notebook prioritises feasible classical-model comparisons rather than exhaustive automated tuning.

## 5.4 Why introduce class weighting?

The initial models tended to prioritise N because N dominates the training examples. Class weighting increases the training penalty for errors on less common classes. The notebook sets class_weight="balanced" for the scikit-learn LR, SVM and RF classifiers. For AdaBoost/XGBoost, it calculates balanced per-example sample weights from the training labels. No synthetic oversampling is needed for this experiment.

## 5.5 Class-weighted validation results

| Model | Accuracy | Balanced accuracy | Macro-F1 | Interpretation |
| --- | --- | --- | --- | --- |
| Linear SVM (C 0.01) | 93.12% | 0.5544 | 0.4914 | Best macro-F1 of the three weighted classical models; S recall rose to 0.54, but F remains difficult. |
| Logistic Regression (C 0.01) | 79.97% | 0.6589 | 0.4617 | Highest balanced accuracy, but much lower overall accuracy and lower macro-F1; more minority predictions trade off against precision. |
| Random Forest {'classifier__max_depth': 15, 'classifier__max_features': 'sqrt', 'classifier__min_samples_leaf': 5, 'classifier__n_estimators': 50} | 96.27% | 0.4396 | 0.4608 | Highest accuracy, but weaker class balance; S recall was about 0.01 and F recall was 0 in the reported weighted run. |

KEY VIVA POINT SVM was not the clear winner on every individual metric. It had the best macro-F1 among the class-weighted classical models; LR had the highest balanced accuracy and RF had the highest accuracy. The primary objective determines which trade-off is preferred. “We selected Linear SVM because it achieved the highest Macro-F1 among the three class-weighted classical models. Random Forest had higher accuracy but poor minority-class recall, while Logistic Regression had higher balanced accuracy but lower Macro-F1. We then evaluated ensemble combinations and selected SVM + XGBoost + AdaBoost because it achieved the highest validation Macro-F1 among the compared ensembles. We acknowledge that the margin over the classical ensemble was modest.” One important distinction: SVM wasn't the winner of every individual experiment. It was the strongest of these three class-weighted classical models by Macro-F1, and the three-model boosting-based combination won the ensemble comparison by the same primary metric.

# 6. Why add XGBoost and AdaBoost?

## 6.1 Why go beyond the first three models?

The three classical models exposed different trade-offs. The next experiment introduced two boosting approaches to check whether sequential tree learners could contribute useful decision patterns that the linear SVM or Random Forest alone did not provide. Their inclusion was an empirical hypothesis, not an assumption that boosting must win.

| Model | Recorded setup | Reason for testing |
| --- | --- | --- |
| AdaBoost | 100 estimators; decision-stump base estimator (tree max_depth=1); learning_rate 0.5; random_state 42; balanced sample weights. | A lightweight sequential boosting method that increases attention to difficult training examples. |
| XGBoost | Multi-class soft-probability objective; 150 trees; max_depth=5; learning_rate 0.05; subsample 0.8; colsample_bytree 0.8; min_child_weight 2; reg_lambda 1; hist tree method; random_state 42; balanced sample weights. | Gradient-boosted trees can model nonlinear interactions among waveform and timing features; subsampling and regularisation constrain complexity. |

## 6.2 Individual validation comparison after class weighting

| Model | Accuracy | Balanced accuracy | Macro-F1 |
| --- | --- | --- | --- |
| Majority vote: LR + SVM + RF | 94.90% | 0.5992 | 0.5332 |
| Linear SVM (class-weighted) | 93.12% | 0.5544 | 0.4914 |
| XGBoost | 93.45% | 0.5297 | 0.4908 |
| AdaBoost | 89.19% | 0.4679 | 0.4750 |
| Logistic Regression (class-weighted) | 79.97% | 0.6589 | 0.4617 |
| Random Forest (class-weighted) | 96.27% | 0.4396 | 0.4608 |

## 6.3 Is using two boosting algorithms weird?

No. Two boosting methods in one ensemble are not inherently invalid: AdaBoost and XGBoost use different learning procedures and may make different errors. The relevant question is whether adding each model improves the validation combination. The notebook tests that question empirically.

XGBoost brings gradient-based boosting with explicit tree regularisation/subsampling controls.

AdaBoost uses a shallow decision tree and reweights examples across boosting rounds.

The ensemble also includes a linear SVM, so the final vote mixes a linear-margin method with two tree-boosting methods.

The limitation is that the notebook does not report an explicit error-correlation or diversity analysis. We can say the combination scored best among the configurations tested; we should not claim the two boosters are guaranteed to be complementary on unseen data.

# 7. Voting ensemble selection: why the three-model winner?

## 7.1 What voting does

Each selected model predicts N, S, V or F for a heartbeat. Hard voting counts the predicted labels and selects the class with the most votes. The chosen procedure is deterministic: when the top vote is tied, use the SVM prediction if SVM is among the tied leaders; otherwise use the first tied label in the saved class order N, S, V, F.

## 7.2 Original complete-run comparison

| Validation combination | No. of models | Accuracy | Balanced accuracy | Macro-F1 | Decision |
| --- | --- | --- | --- | --- | --- |
| SVM + XGBoost + AdaBoost | 3 | 96.43% | 0.5812 | 0.5460 | Selected |
| All five models | 5 | 96.39% | 0.5741 | 0.5445 | Runner-up |
| SVM + RF + XGBoost + AdaBoost | 4 | 96.68% | 0.5310 | 0.5436 | Not selected |
| SVM + RF + LR | 3 | 94.90% | 0.5992 | 0.5332 | Not selected |
| SVM + RF + XGBoost | 3 | 96.43% | 0.5524 | 0.5266 | Not selected |

## 7.3 The honest rationale

1.  Macro-F1 was the declared selection metric, so the combination with the highest validation macro-F1 among the five predefined voting configurations was chosen.

2.  SVM + XGBoost + AdaBoost scored 0.5460, compared with 0.5445 for all five models and 0.5332 for SVM + RF + LR.

3.  The three-model winner is simpler than the five-model ensemble while achieving the highest macro-F1 in this comparison.

4.  The margin over all five models is only 0.0015, so this is a narrow empirical win, not evidence of statistically significant superiority.

DO NOT OVERCLAIM The table compares five predefined combinations; it does not search every possible subset. For example, SVM + LR + XGBoost and SVM + LR + AdaBoost are not in this table. The defensible claim is “best among the combinations tested,” not “the globally optimal ensemble.”

## 7.4 Why not choose SVM + RF + LR?

We did evaluate it. It had higher balanced accuracy (0.5992) than the selected ensemble (0.5812), but lower macro-F1 (0.5332 vs 0.5460). Since macro-F1 was the primary selection metric, it was not the winner. If the project objective changed to prioritising mean class recall, the balanced-accuracy trade-off would deserve further investigation.

## 7.5 Why not just use SVM?

SVM was the best macro-F1 among the class-weighted classical individuals, but its validation macro-F1 was 0.4914, below the selected voting result of 0.5460. Voting can correct some individual errors when the models disagree. However, no claim is made that every vote fixes an error; the improvement is demonstrated by the combined validation score.

# 8. Why stacking was not selected; final test and saving

## 8.1 Voting vs stacking

Stacking trains a second-level Logistic Regression model to combine out-of-fold predictions from SVM, XGBoost and AdaBoost. The notebook creates four one-hot indicator columns per base classifier, giving 12 meta-features. Group-aware three-fold cross-validation on training data creates out-of-fold predictions, so each meta-training prediction is generated by a base model that did not train on that example’s subject group.

| Candidate | Validation accuracy | Balanced accuracy | Macro-F1 | Outcome |
| --- | --- | --- | --- | --- |
| Voting: SVM + XGBoost + AdaBoost | 96.43% | 0.5812 | 0.5460 | Selected |
| Stacking, balanced meta-classifier | 15.67% | 0.4051 | 0.3087 | Not selected |
| Stacking, unweighted meta-classifier | 95.98% | 0.4329 | 0.4511 | Not selected |

The balanced meta-classifier over-corrected: its reported normal-beat recall was only 0.08. The unweighted stack predicted N very strongly and missed S and F in the validation report. Both stacked variants scored below voting by macro-F1, so the simpler hard-vote ensemble was retained.

## 8.2 Final test evaluation of the selected voting ensemble

| Metric | Held-out test result | Meaning |
| --- | --- | --- |
| Accuracy | 86.32% | Overall correctness across 23,521 test examples. |
| Balanced accuracy | 0.5687 | Mean recall across N/S/V/F; class-wise detection is moderate. |
| Macro-F1 | 0.5161 | Equal-class summary; the reported primary test metric. |

| Class | Precision | Recall | F1 | Support |
| --- | --- | --- | --- | --- |
| N | 0.96 | 0.89 | 0.93 | 19,580 |
| S | 0.32 | 0.14 | 0.19 | 592 |
| V | 0.62 | 0.87 | 0.73 | 2,974 |
| F | 0.16 | 0.38 | 0.22 | 375 |

Interpretation: the model is much better at N and V than at S and F. The F recall of 0.38 means it found some fusion examples, but precision of 0.16 indicates many predicted F beats were false positives. S recall of 0.14 remains weak. The final score is an honest limitation to discuss, not a number to hide.

## 8.3 What the model bundle saves

Project/model metadata and the selected validation macro-F1 (0.5460).

The exact class order N, S, V, F and the raw-symbol-to-target mapping.

The 226 feature names and expected feature count, making feature-order checks possible.

The fitted SVM pipeline, XGBoost model, AdaBoost model, and the fitted median imputer used by the boosted models.

Preprocessing settings (360 Hz, filter type/order/cutoffs, zero-phase filtering, window length, and excluded records) and voting/tie-break configuration.

The notebook reloads the artifact and verifies that predictions on a 100-example smoke-test sample match the pre-save predictions. This checks persistence consistency; it is not a substitute for the held-out test evaluation.

# 9. Limitations and reproducibility notes

## 9.1 Limitations to admit confidently

Severe class imbalance: the extracted dataset is 89.48% N; validation has only 25 F examples. F-class metrics are therefore unstable.

The selected ensemble’s advantage over the all-five vote is only 0.0015 macro-F1; a different split or seed may change the ranking.

Only five predefined voting combinations were compared; the search is not exhaustive and no repeated-seed confidence intervals or statistical significance tests are reported.

S and F performance remains inadequate for any clinical use. The prototype is educational/research-oriented and not a diagnostic system.

The classifier only predicts N, S, V or F. It has no explicit unknown/reject option for excluded Q or out-of-scope beats.

Training segments are centred on expert annotation points; inference uses automatic beat detection. Errors in detection/positioning and differences in RR calculation can affect end-to-end features.

Filtering was selected after visual inspection, but the notebook does not report a controlled ablation demonstrating how much filtering improved model metrics.

## 9.3 If asked whether the model is “good”

Answer in terms of the experiment objective: “It is a useful prototype and shows how to build an end-to-end classical-ML workflow. The test results show useful N and V recognition but weaker S and F performance. The project therefore demonstrates experimentation and evaluation, not clinical readiness. Improving minority-class reliability and testing across more subjects would be necessary next steps.”

## 9.4 Next experiments that would be scientifically useful

1.  Make data retrieval fail loudly or retry, then require the expected 44/44 eligible recordings before training.

2.  Save the exact data manifest, split assignments, random seeds, package versions and selected metrics alongside the model bundle.

3.  Compare all relevant three-model subsets or use a predeclared selection plan, still selecting only on validation data.

4.  Run repeated group-aware splits or additional independent subject cohorts to measure ranking variability.

5.  Investigate minority-class thresholding/resampling/calibration carefully, and report per-class precision and recall rather than only accuracy.

6.  Evaluate beat detection and four-class classification as separate stages of the end-to-end inference pipeline.

# 10. Decision register: selected and non-selected models

| Method | Decision in this experiment | Defensible reason |
| --- | --- | --- |
| Logistic Regression | Tested baseline, tuned, class-weighted and used in votes; not in final selected trio. | Useful linear baseline. Weighted LR had best balanced accuracy among weighted individuals (0.6589) but lower macro-F1 (0.4617) than weighted SVM; votes including it scored below selected trio. |
| Linear SVM | Selected in final trio. | Best macro-F1 among weighted classical individuals (0.4914); scaled features; complements tree-based models. Not best on every metric. |
| Random Forest | Tested baseline, tuned, weighted and included in alternative votes; not in final trio. | Tuned baseline macro-F1 0.4590; weighted RF accuracy was high (0.9627) but balanced accuracy 0.4396 and minority recall was weak. Selected votes containing RF did not beat the winner by macro-F1. |
| AdaBoost | Selected in final trio. | Provides a boosting method based on shallow trees and sample weights; contributes to the best observed validation vote, though standalone F performance is weak. |
| XGBoost | Selected in final trio. | Gradient-boosted trees with regularisation/subsampling; individual validation macro-F1 0.4908, and its vote with SVM/AdaBoost scored best among tested combinations. |
| SVM + RF + LR vote | Tested, not selected. | Validation macro-F1 0.5332; balanced accuracy 0.5992. It would be a candidate if balanced accuracy were prioritised over macro-F1. |
| All-five vote | Tested, not selected. | Very close runner-up at 0.5445 macro-F1; selected trio is smaller and scored 0.5460. |
| Four-model SVM + RF + XGB + Ada | Tested, not selected. | Macro-F1 0.5436, below the three-model winner; higher accuracy did not translate to highest macro-F1. |
| SVM + RF + XGBoost | Tested, not selected. | Macro-F1 0.5266, below winner. |
| Stacking: balanced meta-model | Tested, not selected. | Macro-F1 0.3087 and highly distorted class prediction behaviour in validation. |
| Stacking: unweighted meta-model | Tested, not selected. | Macro-F1 0.4511; it failed to identify S and F in the reported validation classification report. |

BEST ONE-LINE DEFENCE “The three-model ensemble was selected because it achieved the top validation macro-F1 among the predefined vote combinations, while using fewer models than the five-model vote. Its margin was small, so I describe it as the best observed configuration on this split, not as universally superior.”

# 11. Likely viva questions and defensible answers (Part I)

Q. What exactly is the model predicting?

A. One of four target labels—N, S, V or F—for an ECG heartbeat window. It does not diagnose a patient or assign every possible MIT-BIH annotation symbol.

Q. Why did you choose MIT-BIH?

A. It provides ECG signals paired with expert beat annotations, which lets me generate labelled heartbeat examples and evaluate predictions against known labels.

Q. Why are there only four classes?

A. I scoped the first experiment to N/S/V/F. Q represents paced or unclassifiable beats in the mapping and was excluded; other event symbols outside the mapping were not model targets.

Q. Why exclude records 102, 104, 107 and 217?

A. They are the four paced recordings excluded under the selected protocol. I preserve that exclusion consistently across extraction and artifact metadata.

Q. Why choose MLII?

A. The notebook prefers MLII when available and used it across the 44 eligible recordings, giving one consistent single-lead input representation.

Q. Why apply a 0.5–40 Hz filter?

A. The notebook compared raw and filtered signals and adopted a Butterworth band-pass to reduce slow baseline drift and high-frequency components while retaining visible ECG peaks. A separate metric ablation was not performed, so I do not claim a quantified accuracy improvement from filtering.

Q. Why use zero-phase filtering?

A. Forward-backward filtering with `sosfiltfilt` avoids a phase shift that could displace waveform features relative to annotated beat locations; second-order sections are a numerically stable filter representation.

Q. Why 0.2 seconds before and 0.4 seconds after the annotation?

A. At 360 Hz this creates 72 pre-annotation samples and 144 post-annotation samples, 216 values total. The fixed window captures a local morphology context and produces the same input width for every example.

Q. Why retain raw waveform samples if you also have summary features?

A. The waveform preserves detailed local shape. Summary amplitude/QRS features provide complementary compact descriptors, while RR features add temporal context.

Q. What are the 226 features?

A. 216 waveform values, six general amplitude statistics, two QRS-slice statistics, and two RR intervals. The six amplitude statistics are mean, standard deviation, minimum, maximum, peak-to-peak and RMS.

Q. Why are RR values sometimes missing?

A. The first or last detected/annotated beat in a record may not have a previous or next neighbour. Those boundary missing values are expected and are imputed with a median fitted on training data only.

Q. How did you avoid patient leakage?

A. I grouped records by subject and used group-aware splitting so the same subject group cannot appear in both training and validation/test. Records 201 and 202 are kept together.

Q. Why intentionally assign fusion-rich subjects?

A. F is rare, so an ordinary group split can leave too few examples to evaluate it. I kept subject groups separate while deliberately placing fusion-rich subjects across partitions. That improves class evaluability but makes the split less population-representative, which I disclose.

Q. Why not use accuracy as the selection metric?

A. Because 89.48% of examples are N. Accuracy can look high while S/F are missed. Macro-F1 gives each class equal weight, so it aligns better with the four-class objective.

# 12. Likely viva questions and defensible answers (Part II)

Q. Was SVM the best individual model?

A. It was the best by macro-F1 among the class-weighted classical individuals at 0.4914. It was not best on every metric: weighted LR had higher balanced accuracy (0.6589), and weighted RF had higher accuracy (96.27%).

Q. Why test class weighting?

A. The target classes are imbalanced. Class weighting or sample weighting increases the training influence of rarer classes without creating synthetic beats. It improved SVM macro-F1 from 0.4125 in the tuned baseline to 0.4914 in the weighted model.

Q. Why add both XGBoost and AdaBoost?

A. They are different boosting procedures, and they might make complementary errors. I tested them individually and in predefined voting combinations. Their inclusion in the final trio is supported by the combination’s validation score, not by a blanket claim that two boosters are always better.

Q. Why not choose SVM + RF + Logistic Regression?

A. It was tested and got macro-F1 0.5332. It had higher balanced accuracy than the selected trio, but macro-F1 was the declared selection criterion and the SVM + XGBoost + AdaBoost vote scored 0.5460.

Q. Why not use all five models if they are more diverse?

A. The all-five vote scored 0.5445 macro-F1, just below the selected trio at 0.5460. The three-model vote is simpler and achieved the highest score among the five predefined combinations. The difference is small and may not persist on other splits.

Q. Why hard voting instead of soft voting?

A. The experiment implements hard voting over predicted classes. This keeps the rule simple and avoids relying on comparable probability estimates across models—LinearSVC in particular does not directly provide calibrated class probabilities in this setup. Soft voting/calibration was not evaluated, so I do not claim hard voting is universally better.

Q. What happens in case of no clear majority?

A. In case of such a tie the SVM which is best between the individual models according to its F1 score is chosen as the deciding model.

Q. Why was stacking not selected?

A. The balanced stacking meta-classifier scored 0.3087 macro-F1; the unweighted version scored 0.4511. Both were below voting, and their class-wise reports showed problematic behaviour. Therefore, stacking added complexity without improving the selected validation metric in this experiment.

Q. Why does validation accuracy look high but test accuracy is lower?

A. The validation and test sets contain different held-out subjects and different class distributions. Validation contains 25 F examples and test contains 375, and the test group mix differs. Accuracy is also dominated by N. Macro-F1 and per-class metrics reveal the weaknesses hidden by accuracy.

Q. What does the final test result mean?

A. The chosen vote achieves 86.32% accuracy, 0.5687 balanced accuracy and 0.5161 macro-F1 on 23,521 held-out examples. It has strong N performance, better V recall, but weak S recall (0.14) and limited F precision (0.16) / recall (0.38).

Q. Can you call it clinically useful?

A. No. The class-level results are limited, especially for S and F, and this experiment does not establish clinical validity. It is an educational prototype for exploring the end-to-end ML workflow.

Q. What does the smoke test prove?

A. It proves that the saved artifact can be reloaded and produces the same predictions as before saving for a 100-example test sample. It checks persistence integrity, not generalisation or clinical performance.

Q. What changed in the later rerun?

A. The later rerun failed to download record 202 (a 502 error), changing the extracted dataset and split composition. The validation winner changed, but the final evaluation/saving cells still referenced the original three-model selection. I therefore keep the original complete run as the reference and do not combine metrics across runs.

Q. What would you improve next?

A. Make data acquisition fail loudly/retry; record package versions and a data manifest; test more group splits; compare ensemble subsets under a predeclared selection rule; investigate minority-class strategies; and evaluate automatic beat detection separately from classification.

# 13. Final cheat sheet: numbers to remember

| Topic | Reference-run value |
| --- | --- |
| Eligible recordings | 44 / 44 processed; paced records 102, 104, 107, 217 excluded. |
| Extracted heartbeat windows | 100,685 |
| Class counts | N 90,094; S 2,781; V 7,008; F 802 |
| Feature vector | 226 values = 216 waveform + 8 summary/QRS + 2 RR |
| Split sizes | Train 58,295; validation 18,869; test 23,521 |
| Primary selection metric | Validation macro-F1 |
| Best voting combination | Class-weighted Linear SVM + XGBoost + AdaBoost |
| Selected validation metrics | Accuracy 96.43%; balanced accuracy 0.5812; macro-F1 0.5460 |
| Runner-up: all five | Macro-F1 0.5445; margin 0.0015 |
| SVM + RF + LR | Macro-F1 0.5332; balanced accuracy 0.5992 |
| Balanced stacking | Macro-F1 0.3087 |
| Unweighted stacking | Macro-F1 0.4511 |
| Final held-out test | Accuracy 86.32%; balanced accuracy 0.5687; macro-F1 0.5161 |
| Test recalls | N 0.89; S 0.14; V 0.87; F 0.38 |
| Persisted artifact | `ecg_lens_model_bundle.joblib`; models, imputer, classes, feature order and preprocessing/voting configuration |

## Final closing statement

CLOSE WITH THIS “The original complete run produced a saved bundle and a metric-driven selection process. SVM + XGBoost + AdaBoost achieved the highest validation macro-F1 among the votes tested. The held-out test still shows weak S/F performance, so this is an experimental ML prototype, not a clinically validated classifier.”